PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

SPIFE® 4000

Cat. No. 2310							           Split Beta SPE Procedure

The SPIFE 4000 Split Beta SPE Method is intended for the separation of serum or urine proteins and quantitation of serum proteins by agarose gel electrophoresis using the SPIFE 4000 system.

SUMMARY

Serum contains over one hundred individual proteins, each with a specific set of functions and subject to specific variation in concentration under different pathologic conditions.1

Since the introduction of moving-boundary electrophoresis by Tiselius2 and the subsequent use of zone electrophoresis, serum proteins have been fractionated on the basis of their electrical charge at a particular pH into five classical fractions: albumin, alpha1, alpha2, beta and gamma proteins. Each of these classical electrophoretic zones, with the exception of albumin, normally contains two or more components. The relative proportions of these fractions have proven to be useful aids in the diagnosis and prognosis of certain disease states.3-5

PRINCIPLE

Proteins are large molecules composed of covalently linked amino acids. Depending on electron distributions resulting from covalent or ionic bonding of structural subgroups, proteins can be either polar or nonpolar at a given pH. Proteins are separated according to their respective electrical charges on agarose gel using both the electrophoretic and electroendosmotic forces present in the system. The proteins are then stained with a visible stain.

REAGENT

1. 	SPIFE 4000 Split Beta SPE Gel

Ingredients: Each gel contains agarose in a tris-barbital/MOPS buffer with calcium lactate, a stabilizer, and a preservative.

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. The gel contains barbital which, in sufficient quantity, can be toxic.

Preparation for Use: The gels are ready for use as packaged.

Storage and Stability: The gels should be stored at room temperature (15 to 30°C) and are stable until the expiration date indicated on the package. The gels must be stored horizontally in the protective packaging in which they are shipped. DO NOT REFRIGERATE OR FREEZE THE GELS.

Signs of Deterioration: Any of the following conditions may indicate deterioration of the gel: (1) crystalline appearance indicating the agarose has been frozen, (2) cracking and peeling indicating drying of the agarose, (3) bacterial growth indicating contamination, (4) thinning of the gel blocks.

2. 	Acid Blue Stain (Serum Stain)

Ingredients: When dissolved as directed, the stain contains 0.5% (w/v) acid blue stain.

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. DO NOT INGEST.

Preparation for Use: Dissolve the entire contents of the vial of stain in 1 L of 5% acetic acid. Mix thoroughly for 30 minutes. Fill the appropriate stain vat.

Storage and Stability: The dry stain should be stored at 15 to 30°C and is stable until the expiration date indicated on the package. The dissolved stain is stable six months when stored at 15 to 30°C.

Signs of Deterioration: The prepared stain should be a homogeneous mixture free of precipitate. Discard if precipitate forms. The stain must be replaced after processing ten gels to avoid contamination.

3. 	Citric Acid Destain

Ingredients: After dissolution, the destain contains 0.3% (w/v) citric acid.

WARNING: FOR IN-VITRO DIAGNOSTIC USE. DO NOT INGEST - IRRITANT.

Preparation for Use: Pour 11 L of deionized water into the Destain vat. Add the entire package of Destain and mix well until completely dissolved. Pour the entire contents of the Destain Additive bottle into the prepared Destain.

Storage and Stability: Store the Destain at 15 to 30°C. It is stable until the expiration date on the package.

Signs of Deterioration: Discard if solution becomes cloudy.

4. 	Destain Additive

Ingredients: The product is a wetting agent.

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. IRRITANT–DO NOT INGEST.

Preparation for Use: Pour the entire contents of the Destain Additive bottle into the 11 L of prepared Citric Acid Destain.

Storage and Stability: The additive should be stored at 15 to 30°C and is stable until the expiration date indicated on the package.

Signs of Deterioration: The additive should be free of precipitate.

5.	Acid Violet Stain (Urine Stain)

Ingredients: The stain is comprised of Acid Violet stain.

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. DO NOT INGEST.

Preparation for Use: Dissolve the dry stain in 1 L of 10% acetic acid and mix thoroughly. Fill the appropriate stain vat.

Storage and Stability: The dry stain should be stored at 15 to 30°C and is stable until the expiration date indicated on the package. The stain solution is stable six months when stored at 15 to 30°C in a closed container.

Signs of Deterioration: The diluted stain should be a homogeneous mixture free of precipitate. The stain must be replaced after processing ten gels to avoid contamination.

INSTRUMENT

A SPIFE 4000 must be used to apply samples, electrophorese, stain, destain, dry and then scan the gels. Refer to the Operator's Manual for detailed instructions.

SPECIMEN COLLECTION AND HANDLING

Specimen: Fresh serum or urine is the specimen of choice. Use of plasma will cause a fibrinogen band to appear as a distinct narrow band between the beta and gamma fractions.

Storage and Stability: If storage is necessary, samples may be stored covered at 15 to 30°C for 4 days, 2 to 8°C for 2 weeks or -20°C for 6 months.6 Urine samples may be stored covered at 2 to 8°C for up to 72 hours or at -20°C for 1 month.

Urine Sample Preparation: Urine samples may be run diluted, neat or concentrated. Swirl samples to homogenize. Vigorous agitation may denature proteins. Centrifuge desired volume at 2000 x g for 5 minutes. Remove supernatant and concentrate as follows:

Total Protein (mg/dL)	Conc. Factor

<	50	100x

50-100	50x

100-300	25x

300-600	10x

>	600	5x

Interfering Factors:

1.	Hemolysis may cause false elevation in the alpha2 and beta fractions.

2.	Inaccurate results may be obtained on specimens left uncovered, due to evaporation.

PROCEDURE

Materials provided: The following materials needed for the procedure are contained in the SPIFE 4000 Split Beta SPE Kit (Cat. No. 2310). Individual items are not available.

SPIFE 4000 Split Beta SPE Gels (10)

Acid Blue Stain (1 vial)

SPIFE 4000 Blotter C (10)

Citric Acid Destain (1 pkg)

SPIFE 4000 Split Beta SPE Applicator Blades (20)

Destain Additive (28 mL)

Materials provided but not contained in the kit:

ITEM	CAT. NO.

SPIFE 4000 Analyzer	1620, 1621

Gel Block Remover	1115

SPE Normal Control	3424

SPE Abnormal Control	3425

SPIFE 4000 White Sample Trays	2315

SPIFE 4000 Cassettes	1630

SPIFE 4000 Maintenance Blotters	2307

SPIFE 4000 Applicator Blades (for urine)	2319

SPIFE 4000 Gel Staging Lid	2308

Acid Violet Stain (for urine)	552351

SPIFE 4000 Replacement Electrodes	1625

Materials needed but not provided:

0.85% saline

5% acetic acid

10% acetic acid

STEP-BY-STEP METHOD

I.	Sample Preparation

A.	Serum

No specimen preparation is necessary. The SPIFE 4000 automatically dispenses samples into the sample tray.

B. 	Urine

Urine specimens cannot be run on the same gel with serum specimens. The applicator blade for urine (Cat No. 2319) is different and must be ordered separately. The running parameters are also different.

For urine specimen volumes measuring less than 500 µL, 40 µL must be hand-pipetted into the sample tray for application. If there is 500 µL or greater, the specimen can be transferred into a tube and automatically dispensed into the tray.

II.		SPIFE 4000 Preparation

1.	Stack the appropriate number of Disposable Sample Trays into the Sample Tray Holder (one tray per gel).

2.	Place Split Beta SPE Applicator Blades (two per gel) in the Applicator Tray.

3.	Fill the designated bottles with deionized water, 0.85% saline and destain.

4.	Add prepared (acid blue or acid violet) stains to the appropriate stain bottle. The stains must be replaced after processing ten gels to avoid contamination.

5.	Fill the DI Water Surfactant jar with deionized water, and replace the lid and tubing. Ensure that the ends of the tubing are below the water level.

6.	Remove the antisera/water reservoir from the antisera station. Lift the cover, fill the "H2O" well and replace the cover.

7.	Turn on the SPIFE 4000. Wait about 3 minutes after turning on the lower unit. Click the SPIFE 4000 icon on the screen for the instrument to initialize.

8.	Using the prompts, prime the sample handler and the surfactant delivery system according to the instructions in the Operator’s Manual.

9.	Carefully open one end of the pouch, remove the gel from the protective packaging and discard the overlay.

10.	Using a SPIFE 4000 Blotter C, gently blot the entire gel. Discard the blotter.

11.	Hold the gel so that the barcode is at the top. Place the gel into the cassette by holding the gel backing in one hand and gently bending the gel. Slide each end of the gel backing into the slots of the cassette to hold it in place. Align the cut out in the gel backing with the alignment pin in the cassette.

12.	Ensure the gel blocks make good contact with the electrodes to prevent skewed patterns.

13.	Place the cassette with the gel into the humidor and cover the topmost cassette with the Gel Staging Lid. Close the humidor lid to minimize gel dehydration.

14.	Repeat Steps 9-13 for each gel needed.

15.	Load patient sample test tubes into the carousel counterclockwise starting at the space numbered 1. Multiple carousels are loaded counterclockwise.

III.	Electrophoresis Parameters

A. Serum

1. Using the instructions provided in the Operator's Manual, select

-- Test Type: Serum Proteins 
Test Name: Serum Proteins

Check the programmed parameters for each of the following processes.

Sample Preparation

Sample Application

Electrophoresis

Predry

Stain

Destain

Dry

Scan

Sample Preparation	Volume (µL) 	40

Primary Wash Time (mm:ss)	00:02

Primary Wash Cycles	0

Samples per gel	28

Application per sample	1

Dilutions 1-28	Neat

Sample Application	Applicator Load Time (mm:ss)	00:30

Applicator Load Speed	95

Application Rows	2

Row 1 Location (mm from gel edge):	45.0

Apply Time (mm:ss)	01:00

Apply Cycles	1

Absorption Time (mm:ss)	00:45

Inter-Gel Start Delay (mm:ss)	00:00

Electrophoresis	Voltage	400

Minimum Current (mA)	10

Maximum Current (mA)	100

Temperature (°C)	21

Time (hh:mm:ss)	00:06:00

Pre-Dry	Temperature (°C)	53

Time (hh:mm:ss)	00:12:00

Stain	Stain Type	Acid Blue

Absorption Time	00:04:00

Destain	Cycles	4

Time (hh:mm:ss)	00:02:00

Agitate	Yes

Dry	Temperature (°C)	60

Time (hh:mm:ss)	00:13:00

Scan	Sequence	1-28

Aperture Size	5

Gain Mode	Auto

Smooth	No

Auto Interpretation	No

Image Contrast	0.0

Fraction Detection

Sensitivity	3.1

Force Fractions	No

Combine Split Beta	Yes

2. Proceed to Section IV. Electrophoresis.

B. Urine

1.	Volume of 500 µL or Greater

a. Pipette 500 µL or more into a test tube and place it in the sample carousel.

b. Using the instructions provided in the Operator's Manual, select

-- Test Type: Urine Proteins 
Test Name: Urine Proteins

Check the programmed parameters for each of the following processes.

Sample Preparation

Sample Application

Electrophoresis

Predry

Stain

Destain

Dry

Scan

Sample Preparation	Volume (µL)	40

Primary Wash Time (mm:ss)	00:02

Primary Wash Cycles	1

Samples per gel	28

Application per sample	1

Dilutions 1-28	Neat

Sample Application	Applicator Load Time (mm:ss)	00:30

Applicator Load Speed	115

Application Rows	2

Row 1 Location (mm from gel edge):	45.0

Apply Time (mm:ss)	01:00

Apply Cycles	3

Absorption Time (mm:ss)	00:00

Inter-Gel Start Delay (mm:ss)	01:00

Electrophoresis	Voltage	400

Minimum Current (mA)	0

Maximum Current (mA)	100

Temperature (°C)	19

Time (hh:mm:ss)	00:07:00

Pre-Dry	Temperature (°C)	53

Time (hh:mm:ss)	00:12:00

Stain	Type	Acid Violet

Time	00:02:00

Destain	Cycles	4

Time (hh:mm:ss)	00:02:00

Agitate	No

Dry	Temperature (°C)	60

Time (hh:mm:ss)	00:13:00

Scan	Sequence	1-28

Aperture Size	5

Gain Mode	Auto

Smooth	No

Auto Interpretation	No

Image Contrast	0.0

Fraction Detection

Sensitivity	5.0

Force Fractions	No

Combine Split Beta	No

c.	Proceed to Section IV. Electrophoresis.

2.	Urine Volumes less than 500 µL

a.	Pipette 40 µL of each urine specimen into the Sample Tray.

b.	Place the Sample Tray with specimens into the Sample Tray Holder.

c.	The SPIFE 4000 User Setup menu should be used to create “Add a Test” parameters for small volumes of urine which need to be manually pipetted. The test parameters will be the same as those given in Section III B.1 except that “Sample Preparation” will not be checked.

d.	Proceed with the testing according to steps in Section IV.

IV. Electrophoresis

1.	Click START on the screen and respond to the analyzer prompts. The analyzer will load samples when appropriate, apply samples, electrophorese, stain, destain, dry and scan the gels.

2.	After scanning, the cassette with the gel will be dropped into the cassette receptacle.

3.	Remove the cassette(s) from the receptacle. If gel storage is required, remove and discard the two gel blocks. Clean or wipe the non-gel side. If not, discard the used gels, blotters, applicator blades and sample trays as biohazardous waste.

4.	Cassettes and carbon electrodes should be washed and dried after each use with deionized water. Refer to the Operator’s Manual for instructions.

Stability of End Product: The completed, dried SPIFE 4000 Split Beta SPE Gel is stable for an indefinite period of time.

Quality Control: SPE Normal Control (Cat. No. 3424) and SPE Abnormal Control (Cat. No. 3425) may be used to verify all phases of the procedure and should be used on each gel run. If desired, a control or patient sample may be diluted 1:7 with 0.85% saline (1 part sample + 6 parts saline) and run with urines. Refer to the package insert provided with the control for assay values.

REFERENCE VALUES

The reference range presented was established with the Split Beta SPE System on 50 normal specimens using the SPIFE 4000 Analyzer. These values are presented as a guideline.

% of_Total Protein

Protein Fraction	X ± 2 S.D.

Albumin	48.1	–	59.5

Alpha1	2.3	–	4.9

Alpha2	6.9	–	13.0

Beta		13.8	–	19.7

Gamma	10.1	–	21.9

Each laboratory should perform its own normal range study.

Variations of Expected Values5

Studies show that values are the same for both males and nonpregnant females. (Some differences are seen in pregnant females at term and in women on oral contraceptives.)

Age has some effect on normal levels. Cord blood has decreased total protein, albumin, alpha2 and beta fractions, slightly increased alpha1 and normal or increased gamma fractions (largely of maternal origin). The gamma globulins drop rapidly until about three months of age, while the other fractions have reached adult levels by this time. Adult levels of the gamma globulins are not reached until 16 years of age. The albumin decreases and beta globulin increases after the age of 40.

RESULTS

Figure 1 illustrates the electrophoretic mobilities of the albumin, alpha1, alpha2, beta and gamma protein bands on SPIFE 4000 Split Beta SPE Gel. The fastest moving band, and normally the most prominent, is the albumin band found closest to the anodic edge of the gel. The faint band next to this is alpha1, followed by alpha2 globulin, split beta and gamma globulins.

Figure 1: A SPIFE 4000 Split Beta SPE Gel showing relative position of the bands.

Figure 2: A scan of a SPIFE 4000 Split Beta SPE pattern.

Calculations of the Unknown

The SPIFE 4000 scanner will automatically calculate and print the relative percent and the absolute value of each band when the total protein is entered. Refer to the Operator's Manual provided with the instrument.

INTERPRETATION OF RESULTS5

Results on normal individuals will cover age and sex-related variations and day-to-day biologic variations. Disease states in which abnormal patterns are observed include inflammatory response, rheumatic disease, liver diseases, protein-loss disorders, plasma cell dyscrasias, pregnancy and genetic deficiencies.

Further Testing Required

The serum protein electropherogram or densitometric tracing should be evaluated for abnormalities. If abnormalities are observed, appropriate follow-up studies should be initiated. These may include immunofixation, quantitation of immunoglobulins, bone marrow examination and other appropriate tests.

LIMITATIONS

Since all electrophoretic procedures are nonlinear, it is critical to fill the wells with the recommended volume of undiluted serum to obtain optimal resolution and reproducible results. Noncompliance with the recommended procedure may affect the results.

SPECIFIC PERFORMANCE CHARACTERISTICS
PRECISION

Within Run: A normal control, an abnormal control and a normal patient were run in replicate on a single gel with the following results: N = 28 each

Normal Control

Protein Fraction 	Mean %	SD	CV

Albumin		56.4	1.1	2.0%

Alpha1			4.0	0.4	10.0%

Alpha2			10.0	0.7	7.0%

Beta			16.8	0.2	1.2%

Gamma		12.7	0.6	4.7%

Abnormal Control

Protein Fraction 	Mean %	SD	CV

Albumin		55.5	0.8	1.4%

Alpha1			3.4	0.2	5.9%

Alpha2			8.6	0.3	3.5%

Beta			13.6	0.4	2.9%

Gamma		18.9	0.6	3.2%

Normal Patient

Protein Fraction 	Mean %	SD	CV

Albumin		55.9	1.1	2.0%

Alpha1			3.4	0.3	8.8%

Alpha2			7.9	0.5	6.3%

Beta			15.5	0.4	2.6%

Gamma		17.4	0.5	2.9%

Between-Run: A normal control, an abnormal control and a normal patient sample were run alternately on eight gels.

Normal Control     N = 80

Protein Fraction 	Mean %	SD	CV

Albumin		58.3	1.0	1.7%

Alpha1			3.9	0.3	8.8%

Alpha2			9.3	0.4	4.7%

Beta			16.7	0.4	2.2%

Gamma		11.9	0.6	4.9%

Abnormal Control     N = 72

Protein Fraction 	Mean %	SD	CV

Albumin		54.0	1.0	1.9%

Alpha1			3.6	0.2	5.2%

Alpha2			8.8	0.3	3.8%

Beta			14.0	0.4	3.2%

Gamma		19.7	0.6	3.0%

Normal Patient      N = 72

Protein Fraction 	Mean %	SD	CV

Albumin		55.7	1.0	1.9%

Alpha1			3.4	0.3	7.4%

Alpha2			8.0	0.4	5.2%

Beta			15.5	0.4	2.7%

Gamma		17.4	0.6	3.6%

Correlation

Fifty normal and fifty abnormal serum samples were analyzed in duplicate using the SPIFE 3000 Split Beta SPE system and the SPIFE 4000 Split Beta SPE system.

N = 100

Y = 0.967X + 0.645

R = 0.999

X = SPIFE 3000 Split Beta SPE

Y = SPIFE 4000 Split Beta SPE

BIBLIOGRAPHY

1.	Alper, C.A., Plasma Protein Measurements as a Diagnostic Aid, N. Eng J Med, 291:287-290, 1974.

2.	Tiselius, A., A New Approach for Electrophoretic Analysis of Colloidal Mixtures, Trans Faraday Soc, 33:524, 1937.

3.	Ritzmann, S.E. and Daniels, J.C., Diagnostic Proteinology: Separation and Characterization of Proteins, Qualitative and Quantitative Assays in Laboratory Medicine, Harper and Row, Inc., Hagerstown, 1979.

4.	Tietz, N.W., ed., Textbook of Clinical Chemistry, W.B. Saunders Co., Philadelphia, pg. 579-582, 1986.

5.	Ritzmann, S.E., ed., Protein Abnormalities Vol I: Physiology of Immunoglobulins Diagnostic and Clinical Aspects, Allen R. Liss, Inc., New York, 1982.

6.	Tietz, N.W., ed., Textbook of Clinical Chemistry, 3rd ed., W.B. Saunders Co., Philadelphia, pg. 524, 1995.

For Sales, Technical and Order Information and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be free from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or refund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyer’s exclusive remedies.

In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages.

The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose.

Pro. 18

10/21(9)