PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

ColoScreen®

A Test for Occult Blood

INTENDED USE

ColoScreen is a guaiac slide test for the qualitative detection of fecal occult blood. It is a useful aid in the diagnosis of a number of gastrointestinal disorders and is recommended for use in:

1. Routine physical examinations

2. Routine hospital testing

3. Mass screening for colorectal cancer

# SUMMARY

The detection of occult blood is critical to many gastrointestinal diseases. The presence of occult blood in fecal material may indicate gastrointestinal pathology such as hemorrhoids, diverticulitis, fissures, colitis or colorectal cancer. Fortunately, these conditions can be detected with several diagnostic methodologies available which include testing of stools for occult blood, complete physical examination with digital examination, and proctosigmoidoscopy. Air contrast barium enema and fiberoptic colonoscopy also contribute significantly to the diagnosis of colonic lesions. Unfortunately, only a small percentage of bowel and rectal cancers are found on digital examination and patients with no symptoms of bowel disease do not readily present themselves for procedures such as proctosigmoidoscopy and barium enema.

The American Cancer Society recommends annual fecal occult blood test plus flexible sigmoidoscopy every 5 years for both men and women beginning at age 50.1 A colonoscopy is suggested every 10 years. Colorectal cancer screening should begin earlier and should be done more often if you have increased risk factors such as personal or family history of colorectal disease.1

ColoScreen is a simple, aesthetic, inexpensive test designed for use in the collection and preparation of stool specimens. It overcomes the instability of guaiac solution and the hypersensitivity of benzidine and ortho-tolidine.2

If a positive result is obtained with the test, a follow-up with additional diagnostic tests, as soon as possible, is essential. As with any occult blood test, results with ColoScreen cannot be considered conclusive evidence of the presence or absence of gastrointestinal bleeding or pathology. The test is not intended as a replacement for other diagnostic procedures such as proctosigmoidoscopy examination, barium enema, and X-ray studies.

# PRINCIPLE

ColoScreen is composed of guaiac impregnated paper enclosed in a cardboard frame which permits sample application to one side, and development and interpretation on the reverse side. The process involves placing two specimens, collected from each of three successive evacuations, onto the guaiac paper.

ColoScreen, like all guaiac paper tests for occult blood, is based on the oxidation of phenolic compounds present in the guaiac (i.e. guaiaconic acids) to quinones resulting in production of the blue color.3 Because of its similarity to the prosthetic group of peroxidase, the hematin portion of the hemoglobin molecule can function in a pseudoenzymatic manner, catalyzing the oxidation of guaiac.

When a fecal specimen containing occult blood is applied to the test paper, contact is made between hemoglobin and the guaiac. A pseudoperoxidase reaction will occur upon the addition of the developer solution, with a blue chromagen formed proportionally to the concentration of hemoglobins. The color reaction will occur after thirty seconds.

Hemoglobin + Developer

Hb + 2H202	2H20 + O2

Oxidation of Guaiac

O2 + Guaiac	Oxidized Guaiac

(Colorless)	 (Blue)

The ColoScreen kits include ColoCheck Monitors which provide a quality control system for each test. The ColoCheck Monitors are incorporated into each slide.

REAGENTS

- ColoScreen Slides and Monitors

# ColoScreen Tape and Monitors

Ingredients: ColoScreen Slides and ColoScreen Tape are made of quality controlled

paper impregnated with guaiac resin. ColoCheck Positive Monitor contains an impregnated

substance which will turn blue if product is functioning properly. The ColoCheck Negative

Monitor consists of guaiac impregnated paper.

WARNING: FOR IN-VITRO DIAGNOSTIC USE.

Preparation for Use: The slide and tape are ready for use as packaged.

Storage and Stability: These products should be stored at room temperature (15 to 30°C)

and are stable until the expiration date indicated on each slide and tape box. Do not use

after expiration date. Both items should be protected from heat, humidity, light, fluorescent

light, U.V. radiation, excessive air flow, or volatile chemicals (e.g. iodine or bleach). Do not

refrigerate or freeze.

Signs of Deterioration: Discoloration of the normally light tan paper may occur if exposed

to sunlight, fluorescent or ultraviolet light. Failure of the control system to react as

expected may be indicative of deterioration of the developer or the slide/tape, and test

results should be regarded as invalid.

2.	ColoScreen Developer

Ingredients: ColoScreen Developer contains < 6% hydrogen peroxide and denaturated

ethanol.

WARNING: FOR IN-VITRO DIAGNOSTIC USE.

DANGER: FLAMMABLE. NEVER PIPETTE BY MOUTH. VAPOR HARMFUL. DO NOT

INGEST OR PLACE IN EYES. May cause blindness, or be fatal if swallowed. Keep away

from heat, sparks, or an open flame. Avoid contact with eyes and skin. Should contact

occur, flush the affected area with water and get immediate medical attention.

Preparation for Use: ColoScreen Developer is ready for use as packaged.

Storage and Stability: ColoScreen Developer should be stored tightly capped at 15 to

30°C protected from heat and light. Under these conditions, the developer will remain

stable until the expiration date indicated on the bottle. Do not use after the expiration

date. Do not substitute reagents from other manufacturers.

Signs of Deterioration: Failure of the ColoCheck Monitors to react as expected may be

indicative of deterioration of the developer or the slide/tape, and the test results should be

regarded as invalid.

SPECIMEN COLLECTION AND HANDLING

Patient Preparation:

A.	It is recommended that the patient be placed on a high residue diet starting 2 days before and continuing through the test period.

DIET MAY INCLUDE:

1. 	Meats: Only small amounts of well-cooked chicken, turkey and tuna.

2. 	Vegetables: Generous amounts of both raw and cooked vegetables including lettuce, corn, spinach, carrots and celery. Avoid raw vegetables with high peroxidase activity such as those listed below.

3. 	Fruits: Plenty of fruits, especially prunes and apples.

4. 	Cereals Bran and bran-containing cereals.

5. 	Moderate amounts of peanuts and popcorn daily. If any of the above foods are known to cause discomfort, the patient is instructed to consult his/her physician.

TO BE AVOIDED:

1. 	Meat: Diet should not include any red or rare meat.

2. 	Raw fruits and vegetables containing high peroxidase activity:

Turnip	Cauliflower	Red radishes

Broccoli	Cantaloupe

Horseradish	Parsnip

B.	Alternately, the special diet may be omitted initially with dietary restrictions imposed upon the re-testing of all positive results. However, because gastrointestinal lesions may bleed intermittently and blood in feces is not distributed uniformly, all patients with positive tests regardless of diet, should have follow-up diagnostic procedures done.

C. Other factors which affect the test:

1. 	Medications: For 7 days prior to and during the testing, do not ingest aspirin4 or other anti-inflammatory medicines. For 2 days prior to and during testing, do not use rectal medicines, or tonics or vitamin preparations which contain Vitamin C (ascorbic acid) in excess of 250 mg per day.

2. 	Bleeding hemorrhoids or open cuts on hands.

3. 	Collection of specimen during menstrual cycle.

4. 	Improper specimen collection.

5. 	Other diseases of the gastrointestinal tract such as colitis, gastritis, diverticulitis and bleeding ulcers.

Specimen Handling: Using the applicators provided, obtain a small sample of the stool from the toilet bowl. It is very important that the stool specimen be applied as a very thin smear to the ColoScreen Slides and Tape. Obtain a second sample of the stool, from a different location, in the same manner. Apply a very thin smear to the slide. Allow the smears to air dry.

The slide smears may be prepared and developed immediately or stored up to 12 days prior to development. Care should be taken so that anything coming into contact with the specimen is free of blood. Because of the nonhomogeneity of the stool, it is recommended that the test be performed on three (3) consecutive evacuations, or as close together as possible.5, 6

Patient specimens and all materials in contact with them should be handled as potentially infectious and should be disposed of using proper precautions.

Return the completed slide to your physician or laboratory as instructed. If the slide is returned by mail, use the foil-back envelope provided. DO NOT use a standard paper envelope, as they are not approved by U.S. Postal Regulations.

Interfering Substances: Ingestion of ascorbic acid (Vitamin C) in high doses has been shown to cause false negative results, and intake should be discontinued 2 days prior to, and during, the test period.7 Peroxidase from fruit and vegetables can cause false positive results.8-11 Elimination of red meat from the diet during the test period eliminates the source of hemoglobin which can cause false positives. Oral medications (such as aspirin, indomethacin, reserpine, phenalbutazone, corticosteroids, etc.) and heavy alcohol consumption may cause irritation or bleeding of the gastrointestinal tract and should be discontinued for 7 days prior to and during the test period.

PROCEDURE

Materials Provided:

The following materials are provided by Helena Laboratories for performance of fecal occult blood tests:

ColoScreen Slides with Monitors

ColoScreen Tape with Monitors

ColoScreen Developer

Specimen Applicators

ColoScreen Mailing Envelopes

Kit combinations have been designed to meet individual laboratory demands. Please see the back page of this package insert for kit listings.

METHOD

A. ColoScreen Slide

1. 	Supply all information listed on the front flap of the 	ColoScreen Slide.

2. 	Open the front flap.

3. 	Using the applicator provided, collect a small amount of stool specimen from the toilet, on one end of applicator. Apply a very thin smear in Box A.

4. 	Reuse applicator to obtain a second sample from a different part of the stool specimen. Apply a very thin smear inside Box B. (On subsequent bowel movements, repeat above steps on additional slides.)

5. 	Allow the specimen to air dry, then close the cover.

6. 	Open perforated window on the back of the slide.

7. 	Apply two (2) drops of ColoScreen Developer to the back side of boxes A and B.

8. 	Read results after 30 seconds and within 2 minutes.

9. 	Record the results; any trace of blue color, within or on the outer rim of the specimen, is positive for occult blood.

ColoCheck Monitors

Note: The procedure for developing the sample test must be completed, interpreted

and recorded before proceeding with the development of ColoCheck Monitors.

1. 	To develop ColoCheck Monitors, place one or two drops of ColoScreen Developer between the Positive and Negative Monitor boxes.

2. 	Read the results after 30 seconds and within 2 minutes.

3. 	Positive ColoCheck Monitor should turn blue, but the Negative ColoCheck Monitor should not have any trace of blue.

# B.	ColoScreen Tape

1. 	Tear a segment of tape from the dispenser. Note that Positive and Negative Monitors are printed on one side of the paper segment.

2. 	Apply a very thin smear of fecal material to the tape on the side opposite the Monitors. Do not smear the specimen over the area where the Monitors are printed.

3. 	Turn the tape segment over and apply two (2) drops of ColoScreen Developer to the back side of the application area.

4. 	Read results after 30 seconds and within 2 minutes.

5. 	Any trace of blue color within the smear is positive for occult blood.

6. 	To develop the Monitors, follow the instructions given in the “ColoCheck Monitor” section above.

Stability of End Product: The color reaction is not permanent. Fading may occur after approximately 2 minutes.

# QUALITY CONTROL

ColoCheck Monitors are provided on each CoIoScreen Slide and each tape segment. This    specially treated area provides assurance that the guaiac-impregnated paper and the ColoScreen Developer are reacting according to product specifications. Positive ColoCheck Monitor is an impregnated substance in a base carrier and will turn blue  within 30 seconds after application of ColoScreen Developer if the test system is reacting according to product specifications. Negative ColoCheck Monitor consists of guaiac impregnated paper and will not turn blue upon addition of ColoScreen Developer.

# INTERPRETATION OF RESULTS

Any trace of blue color within the specimen application area is a positive for occult blood, if ColoCheck Monitors react properly. Remember always to develop the test, interpret, and record results before developing the ColoCheck Monitors. Interpretation of the test should not be done by one who is color blind.

# LIMITATIONS

Results obtained with ColoScreen cannot be considered conclusive evidence of the presence or absence of gastrointestinal bleeding or pathology. False negative results may be obtained, since most bleeding occurs intermittently. ColoScreen tests are designed as a preliminary screen and are not intended to replace other diagnostic procedures such as proctosigmoidoscopy, barium enema or X-ray studies. ColoScreen will detect only hemoglobin released upon hemolysis of the red cell. Should whole blood be applied to the test paper, it is necessary to hemolyze the red cells by the addition of a drop of water before adding the developer. Refer to “Interfering Substances” for a further list of limiting substances.

# EXPECTED RESULTS

The guaiac paper tests detect occult blood but they are not diagnostic for disease. Positive occult blood tests may be obtained for reasons which range from red meat in the diet, diverticulitis, hemorrhoids, colitis to colorectal cancer. Patients who have a positive test should verify that they have followed a proper diet prior to specimen collection, and should immediately consult a physician who can perform definitive tests to determine the cause of bleeding. Patients experiencing symptoms such as persistent diarrhea constipation, abdominal pain, visible bleeding, etc., should consult a physician.

ColoScreen will detect 10 mg of hemoglobin per gram of homogenized fecal material.12 Guaiac impregnated paper has been extensively studied.13-15 Clinical studies illustrate that the guaiac impregnated slide tests yield a positive testing range of 3-5% in screening programs. The false positive rate was between 1-2% during controlled conditions (diet, medical supervision, etc.).

# BIBLIOGRAPHY

1.	American Cancer Society, Colon and Rectum Cancer, August 1, 2001.

2.	Irons, G.V. and Kirsner, J.B., Routine Chemical Tests of the Stool for Occult Blood: An Evaluation, Am J of Med Sci, 294:247-260, 1965.

3. 	Kratochvil, J.F. et al., Isolation and Characterization of Alpha Guaiaconic Acid and the Nature of Guaiacum Blue, Phytochem, 10:2529-2531, 1971.

4.	Grossman, M.I. et al., Fecal Occult Blood Loss Produced by Oral and Intravenous Administration of Various Salicylates, Gastroenterology, 40:383-388, 1961.

5. 	Greegor, D.H., Detection of Silent Colon Cancer in Routine Examination, Cancer, 19:330-337, 1969.

6. 	Greegor, D.H., Occult Blood Testing for Detection of Asymptomatic Colon Cancer, Cancer, 28:131-134, 1971.

7. 	Jaffe, R.M., False-Negative Stool Blood Tests Caused by Ingestion of Ascorbic Acid (Vitamin C), Ann Int Med, 83:824, 1975.

8. 	Winawer, S.J. et al., Current Status of Fecal Occult Blood Testing in Screening for Colorectal Cancer, CA-A-Cancer J for Clinicians, 32(2):100-113, 1982.

9. 	Hagihara, P.F., Proceedings of 68th Annual Clinical Congress, Am College Surgeons, 55-57, 1982.

10. 	Bassett, M.L. and Goulston, K.J., False Positive and Negative Hemoccult Reaction on a Normal Diet and Effect of Diet Restriction, Aust NZ J Med, 10:1-4, 1980.

11. 	Winawer, S.J. and Fleisher, M., Sensitivity and Specificity of the Fecal Occult Blood Test for Colorectal Neoplasia, Gastroenterology, 82:986-991, 1982.

12. 	Ebaugh, Jr., F.G., Quantitative Measurement of Gastrointestinal Blood Loss, J of Laboratory and Clinical Medicine, 53:77, 1959

13. 	Ostrow, J.D. et al., Sensitivity and Reproducibility of Chemical Tests for Fecal Occult Blood with an Emphasis on False-Positive Reactions, Amer J Dig Dis, 18:930, 1973.

14. 	Greegor, D.H., A Progress Report – Detection of Colorectal Cancer Using Guaiac Slides, Cancer, 22(6):360, 1972.

15. 	Hastings, J.B., Mass Screening for Colorectal Cancer, Amer J Surg, 127:288, 1974.

## PRODUCTS AVAILABLE

Cat. No.

ColoScreen III Office Pack	5071

Includes:

100 ColoScreen III Slides

6 x 15 mL ColoScreen Developer

300 Applicators

100 Mailing Envelopes

# ColoScreen Lab Pack 	5072

Includes:

100 ColoScreen Single Slides

2 x 15 mL ColoScreen Developer

100 Applicators

# ColoScreen 1000 Pack 	5076

Includes:

1000 ColoScreen Single Slides

20 x 15 mL ColoScreen Developer

1000 Applicators

# ColoScreen Developer  	5077

20 x 15 mL vials

# ColoScreen Developer-2  	5078

50 x 2 mL vials

# ColoScreen Tape 	5079

Includes:

2 ColoScreen Tape Dispensers each

containing a 100 test roll of guaiac paper

and 2 x 15 mL ColoScreen Developer

# ColoScreen III Lab Pack	5082

Includes:

35 ColoScreen III Slides

2 x 15 mL ColoScreen Developer

100 Applicators

# ColoScreen III 1000 Pack	5083

Includes:

350 ColoScreen III Slides

20 x 15 mL ColoScreen Developer

1000 Applicators

100 ColoScreen III Envelopes	1002

For Sales, Technical and Order Information, and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be tree from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or re-fund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyers exclusive remedies.

In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages.

The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose.

Pro. 35

8/03(11)